# tlog2chart Phase 3 Basic Analysis Report

File: 01-Tykon0527-004-2_53300201.txt
Unit Type: Tykon
Unit FW Ver: 34
Unit FPGA Ver: 5
Unit SN: 53300201
Tool: tlog2chart_P3 v1.3.6
Generated: 2026-03-09 22:07:04

# Analysis Settings

- Active threshold: Pfwd > 3.00 W
- Step4 tolerance: ±2.0%  | settle time: 6.0 ms (Tykon only)
- Step5 reflect rule: Pref > 1.00% of Pfwd  | mask first 3.0 ms
- Pulse mode range: Freq 0.01–50.00 kHz, Duty 10.0–90.0%

# Step 1 — Full Scale Overview

Full plot of Pfwd/Pref/Pdel across the entire tlog.

Figure 1. Full-scale power overview.

# Step 2 — Zoom In to last RF event with 100x smaller time span than Step 1, improving resolution around the last power event.

This plot zooms into the time window around the last RF event. The X-axis time span is 100× smaller than Step 1 to improve resolution near the last power activity.

Figure 2. Active-area power plot (idle removed).

# Step 3 — Customize time range analysis (13.5–15.5 s)

Figure 3. User assigned time range RF cycles.

Step 3 cycle metrics (base on RF on/off signal):

| Cycle | Start row() | SetPt (W) | Pfwd steady (W) | Pfwd peak (W) | Overshoot (W) | Overshoot (%) | Rise to 95% (ms) | RF ON (ms) | Idle before (ms) | Mode / Freq / Duty |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 2 | 15.0 | 14.1 | 24.7 | 9.7 | 64.67 | 0.750 | 50.000 | N/A | Pulse, F=0.400kHz, D=15.0% |
| 2 | 210 | 200.0 | 200.0 | 202.4 | 2.4 | 1.20 | 2.250 | 150.000 | 150.000 | CW |
| 3 | 818 | 15.0 | 14.8 | 27.8 | 12.8 | 85.33 | 0.750 | 50.250 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 4 | 1027 | 200.0 | 200.0 | 205.8 | 5.8 | 2.90 | 2.000 | 150.250 | 149.750 | CW |
| 5 | 1636 | 15.0 | 14.0 | 40.7 | 25.7 | 171.33 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 6 | 1845 | 200.0 | 200.0 | 204.8 | 4.8 | 2.40 | 2.000 | 150.250 | 149.750 | CW |
| 7 | 2454 | 15.0 | 14.0 | 40.9 | 25.9 | 172.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 8 | 2663 | 200.0 | 199.9 | 203.0 | 3.0 | 1.50 | 2.000 | 150.250 | 149.750 | CW |
| 9 | 3272 | 15.0 | 14.0 | 41.3 | 26.3 | 175.33 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 10 | 3481 | 200.0 | 200.0 | 204.6 | 4.6 | 2.30 | 2.000 | 150.250 | 149.750 | CW |
| 11 | 4090 | 15.0 | 14.0 | 41.4 | 26.4 | 176.00 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 12 | 4299 | 200.0 | 200.0 | 201.9 | 1.9 | 0.95 | 2.000 | 150.250 | 149.750 | CW |
| 13 | 4908 | 15.0 | 14.0 | 41.4 | 26.4 | 176.00 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 14 | 5117 | 200.0 | 200.0 | 202.6 | 2.6 | 1.30 | 2.250 | 150.250 | 149.750 | CW |
| 15 | 5726 | 15.0 | 14.3 | 41.6 | 26.6 | 177.33 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 16 | 5934 | 200.0 | 200.0 | 202.1 | 2.1 | 1.05 | 2.250 | 150.000 | 150.000 | CW |
| 17 | 6542 | 15.0 | 13.7 | 29.4 | 14.4 | 96.00 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 18 | 6750 | 200.0 | 200.0 | 202.2 | 2.2 | 1.10 | 2.250 | 150.000 | 150.000 | CW |
| 19 | 7358 | 15.0 | 15.2 | 41.7 | 26.7 | 178.00 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 20 | 7566 | 200.0 | 200.0 | 202.8 | 2.8 | 1.40 | 2.000 | 150.000 | 150.000 | CW |
| 21 | 8174 | 15.0 | 13.8 | 41.5 | 26.5 | 176.67 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 22 | 8382 | 200.0 | 200.0 | 202.8 | 2.8 | 1.40 | 2.250 | 150.000 | 150.000 | CW |
| 23 | 8990 | 15.0 | 13.6 | 41.6 | 26.6 | 177.33 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 24 | 9198 | 200.0 | 200.0 | 202.9 | 2.9 | 1.45 | 2.000 | 150.250 | 150.000 | CW |
| 25 | 9807 | 15.0 | 13.9 | 41.7 | 26.7 | 178.00 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 26 | 10016 | 200.0 | 200.0 | 202.7 | 2.7 | 1.35 | 2.000 | 150.250 | 149.750 | CW |
| 27 | 10625 | 15.0 | 15.1 | 41.8 | 26.8 | 178.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 28 | 10834 | 200.0 | 200.0 | 204.5 | 4.5 | 2.25 | 1.250 | 150.250 | 149.750 | CW |
| 29 | 11443 | 15.0 | 13.4 | 41.9 | 26.9 | 179.33 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 30 | 11652 | 200.0 | 200.0 | 203.3 | 3.3 | 1.65 | 2.000 | 150.250 | 149.750 | CW |
| 31 | 12261 | 15.0 | 13.7 | 41.8 | 26.8 | 178.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 32 | 12470 | 200.0 | 200.0 | 203.4 | 3.4 | 1.70 | 2.000 | 150.250 | 149.750 | CW |
| 33 | 13079 | 15.0 | 14.1 | 41.8 | 26.8 | 178.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 34 | 13288 | 200.0 | 200.0 | 202.8 | 2.8 | 1.40 | 2.000 | 150.250 | 149.750 | CW |
| 35 | 13897 | 15.0 | 13.6 | 41.9 | 26.9 | 179.33 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 36 | 14105 | 200.0 | 200.0 | 203.1 | 3.1 | 1.55 | 2.000 | 150.000 | 150.000 | CW |
| 37 | 14713 | 15.0 | 13.6 | 41.8 | 26.8 | 178.67 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 38 | 14921 | 200.0 | 200.0 | 201.9 | 1.9 | 0.95 | 1.750 | 150.000 | 150.000 | CW |
| 39 | 15529 | 15.0 | 13.5 | 42.0 | 27.0 | 180.00 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 40 | 15737 | 200.0 | 200.0 | 204.9 | 4.9 | 2.45 | 2.000 | 150.000 | 150.000 | CW |
| 41 | 16345 | 15.0 | 13.5 | 41.7 | 26.7 | 178.00 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 42 | 16553 | 200.0 | 200.0 | 204.4 | 4.4 | 2.20 | 1.500 | 150.000 | 150.000 | CW |
| 43 | 17161 | 15.0 | 13.9 | 41.9 | 26.9 | 179.33 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 44 | 17369 | 200.0 | 200.0 | 204.1 | 4.1 | 2.05 | 2.000 | 150.250 | 150.000 | CW |
| 45 | 17978 | 15.0 | 13.4 | 41.8 | 26.8 | 178.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 46 | 18187 | 200.0 | 200.0 | 204.2 | 4.2 | 2.10 | 2.000 | 150.250 | 149.750 | CW |
| 47 | 18796 | 15.0 | 13.7 | 41.8 | 26.8 | 178.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 48 | 19005 | 200.0 | 199.9 | 202.8 | 2.8 | 1.40 | 2.250 | 150.250 | 149.750 | CW |
| 49 | 19614 | 15.0 | 14.0 | 35.4 | 20.4 | 136.00 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 50 | 19823 | 200.0 | 200.0 | 203.9 | 3.9 | 1.95 | 2.000 | 150.250 | 149.750 | CW |
| 51 | 20432 | 15.0 | 13.5 | 41.7 | 26.7 | 178.00 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 52 | 20641 | 200.0 | 200.0 | 203.9 | 3.9 | 1.95 | 2.000 | 150.250 | 149.750 | CW |
| 53 | 21250 | 15.0 | 14.2 | 31.0 | 16.0 | 106.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 54 | 21459 | 200.0 | 200.0 | 204.6 | 4.6 | 2.30 | 2.000 | 150.250 | 149.750 | CW |
| 55 | 22068 | 15.0 | 14.0 | 41.9 | 26.9 | 179.33 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 56 | 22276 | 200.0 | 200.0 | 204.3 | 4.3 | 2.15 | 2.000 | 150.000 | 150.000 | CW |
| 57 | 22884 | 15.0 | 13.7 | 41.8 | 26.8 | 178.67 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 58 | 23092 | 200.0 | 200.0 | 204.4 | 4.4 | 2.20 | 2.000 | 150.000 | 150.000 | CW |
| 59 | 23700 | 15.0 | 13.5 | 41.9 | 26.9 | 179.33 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 60 | 23908 | 200.0 | 200.0 | 203.9 | 3.9 | 1.95 | 2.000 | 150.000 | 150.000 | CW |
| 61 | 24516 | 15.0 | 13.4 | 41.9 | 26.9 | 179.33 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 62 | 24724 | 200.0 | 199.9 | 202.2 | 2.2 | 1.10 | 2.000 | 150.000 | 150.000 | CW |
| 63 | 25332 | 15.0 | 13.5 | 41.6 | 26.6 | 177.33 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 64 | 25540 | 200.0 | 199.9 | 202.2 | 2.2 | 1.10 | 2.250 | 150.250 | 150.000 | CW |
| 65 | 26149 | 15.0 | 14.0 | 31.5 | 16.5 | 110.00 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 66 | 26358 | 200.0 | 200.0 | 202.0 | 2.0 | 1.00 | 2.000 | 150.250 | 149.750 | CW |
| 67 | 26967 | 15.0 | 13.4 | 41.9 | 26.9 | 179.33 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 68 | 27176 | 200.0 | 200.0 | 201.6 | 1.6 | 0.80 | 2.000 | 150.250 | 149.750 | CW |
| 69 | 27785 | 15.0 | 13.6 | 41.9 | 26.9 | 179.33 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 70 | 27994 | 200.0 | 199.9 | 203.4 | 3.4 | 1.70 | 1.750 | 150.250 | 149.750 | CW |
| 71 | 28603 | 15.0 | 13.8 | 29.5 | 14.5 | 96.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 72 | 28812 | 200.0 | 199.9 | 202.4 | 2.4 | 1.20 | 2.000 | 150.250 | 149.750 | CW |
| 73 | 29421 | 15.0 | 13.5 | 41.9 | 26.9 | 179.33 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 74 | 29630 | 200.0 | 200.0 | 204.0 | 4.0 | 2.00 | 2.250 | 150.500 | 149.750 | CW |
| 75 | 30240 | 15.0 | 13.7 | 41.9 | 26.9 | 179.33 | 0.750 | 50.000 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 76 | 30448 | 200.0 | 200.0 | 204.3 | 4.3 | 2.15 | 2.000 | 150.250 | 150.000 | CW |
| 77 | 31057 | 15.0 | 13.8 | 41.8 | 26.8 | 178.67 | 0.750 | 50.000 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 78 | 31265 | 200.0 | 200.0 | 204.0 | 4.0 | 2.00 | 2.000 | 150.250 | 150.000 | CW |
| 79 | 31874 | 15.0 | 13.5 | 42.1 | 27.1 | 180.67 | 0.750 | 50.000 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 80 | 32082 | 200.0 | 200.0 | 203.8 | 3.8 | 1.90 | 1.250 | 150.000 | 150.000 | CW |
| 81 | 32690 | 15.0 | 13.4 | 42.0 | 27.0 | 180.00 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 82 | 32898 | 200.0 | 200.0 | 203.6 | 3.6 | 1.80 | 2.250 | 150.000 | 150.000 | CW |
| 83 | 33506 | 15.0 | 13.6 | 41.8 | 26.8 | 178.67 | 0.750 | 50.250 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 84 | 33715 | 200.0 | 200.0 | 204.9 | 4.9 | 2.45 | 2.000 | 150.250 | 149.750 | CW |
| 85 | 34324 | 15.0 | 14.2 | 29.8 | 14.8 | 98.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 86 | 34533 | 200.0 | 200.0 | 204.3 | 4.3 | 2.15 | 2.000 | 150.250 | 149.750 | CW |
| 87 | 35142 | 15.0 | 13.9 | 41.7 | 26.7 | 178.00 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 88 | 35351 | 200.0 | 200.0 | 203.6 | 3.6 | 1.80 | 2.000 | 150.250 | 149.750 | CW |
| 89 | 35960 | 15.0 | 13.8 | 30.6 | 15.6 | 104.00 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 90 | 36169 | 200.0 | 200.0 | 204.2 | 4.2 | 2.10 | 2.250 | 150.250 | 149.750 | CW |
| 91 | 36778 | 15.0 | 13.5 | 41.7 | 26.7 | 178.00 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 92 | 36987 | 200.0 | 200.0 | 208.3 | 8.3 | 4.15 | 1.250 | 150.250 | 149.750 | CW |
| 93 | 37596 | 15.0 | 14.3 | 32.7 | 17.7 | 118.00 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 94 | 37804 | 200.0 | 200.0 | 204.2 | 4.2 | 2.10 | 2.000 | 150.000 | 150.000 | CW |
| 95 | 38412 | 15.0 | 13.4 | 29.5 | 14.5 | 96.67 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 96 | 38620 | 200.0 | 200.0 | 204.3 | 4.3 | 2.15 | 2.000 | 150.000 | 150.000 | CW |
| 97 | 39228 | 15.0 | 13.7 | 41.7 | 26.7 | 178.00 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 98 | 39436 | 200.0 | 200.0 | 203.4 | 3.4 | 1.70 | 2.250 | 150.000 | 150.000 | CW |
| 99 | 40044 | 15.0 | 13.5 | 41.7 | 26.7 | 178.00 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 100 | 40252 | 200.0 | 200.0 | 204.3 | 4.3 | 2.15 | 2.250 | 150.000 | 150.000 | CW |
| 101 | 40860 | 15.0 | 13.4 | 41.7 | 26.7 | 178.00 | 0.750 | 50.250 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 102 | 41069 | 200.0 | 200.0 | 203.7 | 3.7 | 1.85 | 1.250 | 150.250 | 149.750 | CW |
| 103 | 41678 | 15.0 | 13.4 | 41.7 | 26.7 | 178.00 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 104 | 41887 | 200.0 | 199.9 | 203.5 | 3.5 | 1.75 | 1.500 | 150.250 | 149.750 | CW |
| 105 | 42496 | 15.0 | 14.7 | 41.8 | 26.8 | 178.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 106 | 42705 | 200.0 | 200.0 | 204.0 | 4.0 | 2.00 | 2.750 | 150.250 | 149.750 | CW |
| 107 | 43314 | 15.0 | 13.7 | 41.8 | 26.8 | 178.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 108 | 43523 | 200.0 | 200.0 | 205.0 | 5.0 | 2.50 | 2.000 | 150.500 | 149.750 | CW |
| 109 | 44133 | 15.0 | 13.9 | 30.2 | 15.2 | 101.33 | 0.750 | 50.250 | 349.500 | Pulse, F=0.400kHz, D=15.0% |
| 110 | 44342 | 200.0 | 200.0 | 204.5 | 4.5 | 2.25 | 1.750 | 150.250 | 149.750 | CW |
| 111 | 44951 | 15.0 | 14.2 | 29.5 | 14.5 | 96.67 | 0.750 | 50.250 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 112 | 45160 | 200.0 | 199.9 | 203.0 | 3.0 | 1.50 | 2.000 | 150.500 | 149.750 | CW |
| 113 | 45770 | 15.0 | 13.9 | 30.0 | 15.0 | 100.00 | 0.750 | 50.000 | 349.750 | Pulse, F=0.400kHz, D=15.0% |
| 114 | 45978 | 200.0 | 200.0 | 203.9 | 3.9 | 1.95 | 1.750 | 150.000 | 150.000 | CW |
| 115 | 46586 | 15.0 | 13.5 | 30.8 | 15.8 | 105.33 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 116 | 46794 | 200.0 | 200.0 | 204.2 | 4.2 | 2.10 | 1.250 | 150.000 | 150.000 | CW |
| 117 | 47402 | 15.0 | 13.6 | 41.7 | 26.7 | 178.00 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 118 | 47610 | 200.0 | 200.0 | 203.0 | 3.0 | 1.50 | 2.000 | 150.000 | 150.000 | CW |
| 119 | 48218 | 15.0 | 13.5 | 30.4 | 15.4 | 102.67 | 0.750 | 50.000 | 350.000 | Pulse, F=0.400kHz, D=15.0% |
| 120 | 48426 | 200.0 | 200.0 | 204.1 | 4.1 | 2.05 | 2.000 | 150.000 | 150.000 | CW |

Detected RF cycles: 120. The plot shows user assigned time range.

# Step 4 — Forward Power vs Setpoint (Tykon only)

Setpoint column: SetPt. Rule: Pfwd must reach ±2.0% of setpoint within 6.0 ms after RF ON. Steady-state mean after settle window must not be low beyond tolerance.

Flagged cycles: 58

| Cycle | Start row() | Start t(ms) | Setpoint (W) | Steady Pfwd mean (W) | Flags |
| --- | --- | --- | --- | --- | --- |
| 1 | 3 | 633.250 | 15.00 | 14.07 | Low Pfwd vs SetPt (mean 14.07 W, SetPt 15.00 W) |
| 3 | 819 | 1333.250 | 15.00 | 14.85 | Slow settling (> 6.0 ms to reach ±2.0%) |
| 5 | 1637 | 2033.250 | 15.00 | 13.96 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.96 W, SetPt 15.00 W) |
| 7 | 2455 | 2733.250 | 15.00 | 13.98 | Low Pfwd vs SetPt (mean 13.98 W, SetPt 15.00 W) |
| 9 | 3273 | 3433.250 | 15.00 | 13.97 | Low Pfwd vs SetPt (mean 13.97 W, SetPt 15.00 W) |
| 11 | 4091 | 4133.250 | 15.00 | 14.05 | Low Pfwd vs SetPt (mean 14.05 W, SetPt 15.00 W) |
| 13 | 4909 | 4833.250 | 15.00 | 14.03 | Low Pfwd vs SetPt (mean 14.03 W, SetPt 15.00 W) |
| 15 | 5727 | 5533.500 | 15.00 | 14.31 | Low Pfwd vs SetPt (mean 14.31 W, SetPt 15.00 W) |
| 17 | 6543 | 6233.500 | 15.00 | 13.74 | Low Pfwd vs SetPt (mean 13.74 W, SetPt 15.00 W) |
| 21 | 8175 | 7633.500 | 15.00 | 13.84 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.84 W, SetPt 15.00 W) |
| 23 | 8991 | 8333.500 | 15.00 | 13.56 | Low Pfwd vs SetPt (mean 13.56 W, SetPt 15.00 W) |
| 25 | 9808 | 9033.500 | 15.00 | 13.87 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.87 W, SetPt 15.00 W) |
| 29 | 11444 | 10433.500 | 15.00 | 13.37 | Low Pfwd vs SetPt (mean 13.37 W, SetPt 15.00 W) |
| 31 | 12262 | 11133.500 | 15.00 | 13.71 | Low Pfwd vs SetPt (mean 13.71 W, SetPt 15.00 W) |
| 33 | 13080 | 11833.500 | 15.00 | 14.15 | Low Pfwd vs SetPt (mean 14.15 W, SetPt 15.00 W) |
| 35 | 13898 | 12533.750 | 15.00 | 13.60 | Low Pfwd vs SetPt (mean 13.60 W, SetPt 15.00 W) |
| 37 | 14714 | 13233.750 | 15.00 | 13.58 | Low Pfwd vs SetPt (mean 13.58 W, SetPt 15.00 W) |
| 39 | 15530 | 13933.750 | 15.00 | 13.51 | Low Pfwd vs SetPt (mean 13.51 W, SetPt 15.00 W) |
| 41 | 16346 | 14633.750 | 15.00 | 13.50 | Low Pfwd vs SetPt (mean 13.50 W, SetPt 15.00 W) |
| 43 | 17162 | 15333.750 | 15.00 | 13.95 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.95 W, SetPt 15.00 W) |
| 45 | 17979 | 16033.750 | 15.00 | 13.44 | Low Pfwd vs SetPt (mean 13.44 W, SetPt 15.00 W) |
| 47 | 18797 | 16733.750 | 15.00 | 13.66 | Low Pfwd vs SetPt (mean 13.66 W, SetPt 15.00 W) |
| 49 | 19615 | 17433.750 | 15.00 | 14.04 | Low Pfwd vs SetPt (mean 14.04 W, SetPt 15.00 W) |
| 51 | 20433 | 18133.750 | 15.00 | 13.53 | Low Pfwd vs SetPt (mean 13.53 W, SetPt 15.00 W) |
| 53 | 21251 | 18833.750 | 15.00 | 14.18 | Low Pfwd vs SetPt (mean 14.18 W, SetPt 15.00 W) |
| 55 | 22069 | 19534.000 | 15.00 | 13.99 | Low Pfwd vs SetPt (mean 13.99 W, SetPt 15.00 W) |
| 57 | 22885 | 20234.000 | 15.00 | 13.73 | Low Pfwd vs SetPt (mean 13.73 W, SetPt 15.00 W) |
| 59 | 23701 | 20934.000 | 15.00 | 13.52 | Low Pfwd vs SetPt (mean 13.52 W, SetPt 15.00 W) |
| 61 | 24517 | 21634.000 | 15.00 | 13.41 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.41 W, SetPt 15.00 W) |
| 63 | 25333 | 22334.000 | 15.00 | 13.54 | Low Pfwd vs SetPt (mean 13.54 W, SetPt 15.00 W) |
| 65 | 26150 | 23034.000 | 15.00 | 14.03 | Low Pfwd vs SetPt (mean 14.03 W, SetPt 15.00 W) |
| 67 | 26968 | 23734.000 | 15.00 | 13.43 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.43 W, SetPt 15.00 W) |
| 69 | 27786 | 24434.000 | 15.00 | 13.58 | Low Pfwd vs SetPt (mean 13.58 W, SetPt 15.00 W) |
| 71 | 28604 | 25134.000 | 15.00 | 13.86 | Low Pfwd vs SetPt (mean 13.86 W, SetPt 15.00 W) |
| 73 | 29422 | 25834.000 | 15.00 | 13.57 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.57 W, SetPt 15.00 W) |
| 75 | 30241 | 26534.250 | 15.00 | 13.73 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.73 W, SetPt 15.00 W) |
| 77 | 31058 | 27234.250 | 15.00 | 13.81 | Low Pfwd vs SetPt (mean 13.81 W, SetPt 15.00 W) |
| 79 | 31875 | 27934.250 | 15.00 | 13.48 | Low Pfwd vs SetPt (mean 13.48 W, SetPt 15.00 W) |
| 81 | 32691 | 28634.250 | 15.00 | 13.37 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.37 W, SetPt 15.00 W) |
| 83 | 33507 | 29334.250 | 15.00 | 13.56 | Low Pfwd vs SetPt (mean 13.56 W, SetPt 15.00 W) |
| 85 | 34325 | 30034.250 | 15.00 | 14.18 | Low Pfwd vs SetPt (mean 14.18 W, SetPt 15.00 W) |
| 87 | 35143 | 30734.250 | 15.00 | 13.94 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.94 W, SetPt 15.00 W) |
| 89 | 35961 | 31434.250 | 15.00 | 13.87 | Low Pfwd vs SetPt (mean 13.87 W, SetPt 15.00 W) |
| 91 | 36779 | 32134.250 | 15.00 | 13.56 | Low Pfwd vs SetPt (mean 13.56 W, SetPt 15.00 W) |
| 93 | 37597 | 32834.500 | 15.00 | 14.22 | Low Pfwd vs SetPt (mean 14.22 W, SetPt 15.00 W) |
| 95 | 38413 | 33534.500 | 15.00 | 13.42 | Low Pfwd vs SetPt (mean 13.42 W, SetPt 15.00 W) |
| 97 | 39229 | 34234.500 | 15.00 | 13.68 | Low Pfwd vs SetPt (mean 13.68 W, SetPt 15.00 W) |
| 99 | 40045 | 34934.500 | 15.00 | 13.53 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.53 W, SetPt 15.00 W) |
| 101 | 40861 | 35634.500 | 15.00 | 13.44 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.44 W, SetPt 15.00 W) |
| 103 | 41679 | 36334.500 | 15.00 | 13.40 | Slow settling (> 6.0 ms to reach ±2.0%); Low Pfwd vs SetPt (mean 13.40 W, SetPt 15.00 W) |
| 105 | 42497 | 37034.500 | 15.00 | 14.67 | Low Pfwd vs SetPt (mean 14.67 W, SetPt 15.00 W) |
| 107 | 43315 | 37734.500 | 15.00 | 13.77 | Low Pfwd vs SetPt (mean 13.77 W, SetPt 15.00 W) |
| 109 | 44134 | 38434.500 | 15.00 | 13.88 | Low Pfwd vs SetPt (mean 13.88 W, SetPt 15.00 W) |
| 111 | 44952 | 39134.500 | 15.00 | 14.18 | Low Pfwd vs SetPt (mean 14.18 W, SetPt 15.00 W) |
| 113 | 45771 | 39834.750 | 15.00 | 13.88 | Low Pfwd vs SetPt (mean 13.88 W, SetPt 15.00 W) |
| 115 | 46587 | 40534.750 | 15.00 | 13.49 | Low Pfwd vs SetPt (mean 13.49 W, SetPt 15.00 W) |
| 117 | 47403 | 41234.750 | 15.00 | 13.57 | Low Pfwd vs SetPt (mean 13.57 W, SetPt 15.00 W) |
| 119 | 48219 | 41934.750 | 15.00 | 13.55 | Low Pfwd vs SetPt (mean 13.55 W, SetPt 15.00 W) |

Note: table capped at 200 rows for readability.

# Step 5 — Reflected Power (Pref high)

Rule: Pref > 1.00% of Pfwd after masking the first 3.0 ms.

Flagged cycles: 60

| Cycle | Start row() | First bad row() | Worst row() | Worst Pref (W) | Worst Pref/Pfwd (%) |
| --- | --- | --- | --- | --- | --- |
| 1 | 3 | 25 | 175 | 11.10 | 68.10 |
| 3 | 819 | 831 | 1011 | 10.70 | 68.15 |
| 5 | 1637 | 1648 | 1699 | 13.00 | 70.27 |
| 7 | 2455 | 2467 | 2507 | 13.20 | 69.47 |
| 9 | 3273 | 3284 | 3345 | 12.70 | 69.78 |
| 11 | 4091 | 4103 | 4143 | 13.30 | 70.00 |
| 13 | 4909 | 4921 | 4961 | 13.30 | 70.00 |
| 15 | 5727 | 5739 | 5799 | 12.70 | 69.78 |
| 17 | 6543 | 6558 | 6685 | 12.10 | 69.54 |
| 19 | 7359 | 7371 | 7411 | 12.90 | 69.73 |
| 21 | 8175 | 8187 | 8237 | 13.20 | 70.59 |
| 23 | 8991 | 9003 | 9033 | 14.10 | 71.21 |
| 25 | 9808 | 9820 | 9860 | 13.40 | 70.53 |
| 27 | 10626 | 10638 | 10678 | 12.90 | 69.73 |
| 29 | 11444 | 11456 | 11496 | 13.90 | 71.28 |
| 31 | 12262 | 12274 | 12304 | 14.00 | 71.07 |
| 33 | 13080 | 13092 | 13152 | 12.70 | 69.78 |
| 35 | 13898 | 13910 | 13940 | 14.10 | 71.57 |
| 37 | 14714 | 14726 | 14756 | 14.00 | 71.07 |
| 39 | 15530 | 15542 | 15582 | 13.70 | 70.98 |
| 41 | 16346 | 16358 | 16388 | 14.30 | 71.50 |
| 43 | 17162 | 17174 | 17224 | 13.20 | 70.21 |
| 45 | 17979 | 17991 | 18021 | 14.30 | 71.50 |
| 47 | 18797 | 18809 | 18839 | 14.00 | 71.07 |
| 49 | 19615 | 19627 | 19677 | 13.10 | 70.43 |
| 51 | 20433 | 20445 | 20475 | 14.20 | 71.36 |
| 53 | 21251 | 21263 | 21343 | 12.40 | 69.66 |
| 55 | 22069 | 22081 | 22141 | 13.00 | 70.27 |
| 57 | 22885 | 22897 | 22927 | 13.90 | 70.92 |
| 59 | 23701 | 23713 | 23753 | 13.80 | 71.13 |
| 61 | 24517 | 24529 | 24559 | 13.80 | 70.77 |
| 63 | 25333 | 25345 | 25375 | 14.10 | 71.21 |
| 65 | 26150 | 26162 | 26202 | 13.20 | 70.21 |
| 67 | 26968 | 26980 | 27010 | 14.00 | 71.43 |
| 69 | 27786 | 27798 | 27828 | 14.20 | 71.36 |
| 71 | 28604 | 28620 | 28676 | 12.80 | 69.95 |
| 73 | 29422 | 29434 | 29474 | 13.90 | 71.28 |
| 75 | 30241 | 30253 | 30273 | 14.30 | 70.79 |
| 77 | 31058 | 31070 | 31090 | 14.50 | 70.73 |
| 79 | 31875 | 31887 | 31917 | 14.30 | 71.50 |
| 81 | 32691 | 32703 | 32733 | 13.90 | 71.28 |
| 83 | 33507 | 33519 | 33539 | 14.50 | 71.43 |
| 85 | 34325 | 34337 | 34397 | 12.90 | 70.49 |
| 87 | 35143 | 35155 | 35195 | 13.10 | 70.43 |
| 89 | 35961 | 35973 | 36033 | 13.30 | 70.74 |
| 91 | 36779 | 36791 | 36821 | 14.20 | 71.36 |
| 93 | 37597 | 37609 | 37649 | 13.30 | 70.74 |
| 95 | 38413 | 38425 | 38465 | 13.80 | 71.13 |
| 97 | 39229 | 39241 | 39291 | 13.40 | 70.90 |
| 99 | 40045 | 40057 | 40097 | 13.60 | 70.83 |
| 101 | 40861 | 40873 | 40913 | 14.20 | 71.72 |
| 103 | 41679 | 41691 | 41711 | 14.60 | 71.92 |
| 105 | 42497 | 42509 | 42539 | 13.00 | 70.27 |
| 107 | 43315 | 43327 | 43347 | 14.30 | 71.50 |
| 109 | 44134 | 44146 | 44216 | 13.10 | 70.43 |
| 111 | 44952 | 44964 | 45024 | 13.00 | 70.27 |
| 113 | 45771 | 45783 | 45823 | 13.50 | 70.68 |
| 115 | 46587 | 46599 | 46629 | 14.20 | 71.72 |
| 117 | 47403 | 47415 | 47455 | 13.60 | 70.83 |
| 119 | 48219 | 48232 | 48281 | 13.60 | 70.83 |

Note: table capped at 200 rows.

# Pulse Mode Range Alarms

Pulse-mode rows: 10712

Alarms: 0